ASPARTATE AMINOTRANSFERASE (AST) PROCEDURE FOR ALINITY

PRINCIPLE

AST is found in most body fluids, but not in urine except in instances of kidney lesions. The greatest concentrations of ASP are found in heart, liver, muscle, and kidney tissues. Following myocardial infarction, AST in serum begins to increase within 6-8 hours of onset of pain, reaching a peak within 18-24 hours and falling to normal by the fourth or fifth day. Serum values may increase to 10-15 times normal levels and the increase is roughly proportional to the degree of tissue damage.

Methodology: NADH (without P-5’ -P)

SCOPE

The AST assay is used for the quantitation of aspartate aminotransferase in human serum or plasma on the Alinity C analyzer performed in the Texas Health Allen laboratory.

SAFETY

Bloodborne Pathogen Risk – high

Chemical Risk – low

PPE Required – lab coats, gloves, splash shields (goggles, face shields, counter-mounted shields)

- Do not use components beyond the expiration date.

- Do not mix materials from different kit lot numbers.

- This product requires the handling of human specimens. It is recommended that all human sourced materials are considered potentially infectious and be handled in accordance with OSHA Standard on Bloodborne Pathogens.

SPECIMEN REQUIRMENTS

Serum and plasma are acceptable specimens.

Serum: Use serum with or without gel barrier collected by standard venipuncture techniques in plastic tubes. Ensure complete clot formation has taken place prior to centrifugation.  Separate from red blood cells as soon after collection as possible.

Plasma: Use plasma with or without gel barrier. Separate from red blood cells as soon after collection as possible.

SPECIMEN STORAGE STABILITY

30C is good for 4 days

2-8C is good for 7 days

-20C is good for 12 weeks

Note: Allow specimens to reach room temperature prior to testing. Stored specimens must be mixed prior to testing and recentrifuged if they contain particulate matter.

REAGENT HANDLING

- Upon receipt, place reagent cartridge in an upright position for at least 1 hour before use to allow bubbles that may have formed to dissipate.

- If a reagent cartridge is dropped, place in an upright position for 1 hour before use.

- Reagents are susceptible to the formation of foam and bubbles. Bubbles may interfere with the detection of the reagent level in the cartridge and cause insufficient reagent aspiration that may adversely affect results.

REAGENT STORAGE

- Unopened stored at 2-8C are good until the original expiration date.

- Onboard reagent stability is 30 days.

Note: Reagent on board stability is determined by the instrument when the reagent is put into use and stored electronically. The instrument flags the user with information about expiration dates and will not perform testing beyond the reagent’s on board expiration date.

No reagents will be used beyond the expiration date.

MAINTENANCE

Perform daily or as needed maintenance. Refer to Alinity System Operations Manual. Daily, weekly, monthly, and quarterly maintenance operations and instructions are described in the online manual.

CALIBRATION

Calibration is stable for approximately 30 days (720 hours) and calibration is required with each new lot number of reagent. Verify calibration with at least two levels of controls. If control results are not within acceptable ranges, recalibration may be necessary.

To order a calibration select CREATE ORDER from the home screen, then select CALIBRATION from the tabs along the top. Enter the rack number in the rack field and the position number in the position field. In the Assays section, select the assays to be ordered and select ADD ORDER tab along the bottom of the screen.

Calibration may be necessary in the following conditions:

- Out of range control values

- Significant shift or trend in Quality Control (QC) values

- A bulk solution is placed into service

- Any dispense system component is replaced

- Any c-side maintenance procedure is performed

CALIBRATION VERIFICATION

Calibration is validated by running 2 levels of QC after each calibration. If QC values are acceptable, the calibration is considered valid, and patients can be run.

QUALITY CONTROL

Material: BioRad Unassayed Chemistry Control, levels 1 and 2.

Storage and Stability: Stable until the expiration date printed on the bottle when stored at -20C. Once thawed, the control is stable for 15 days at 2-8C, except for bilirubin which is stable for 6 days and AST which is stable for 10 days. Thaw the control bottle at room temperature (18-25C) until completely thawed. DO NOT THAW CONTROLS IN THE FRIDGE. Swirl gently to ensure homogeneity. Use immediately or refrigerate.

Frequency: 2 levels of QC are to be run every 24 hours.

Interpretation and Troubleshooting: If QC results do not meet the acceptance criteria, patient values may be suspect. Review QC results and acceptance criteria following a change of reagent or calibrator lot.

PROCEDURE

- Ensure that sufficient reagent is on board and that it is calibrated.

- Verify sample integrity, sufficient volume, and approved sample type.

- Load sample into appropriate rack and lane on the Alinity.

RESULT REPORTING

AMR:

CRR:

Critical:

If a value exceeds the low end of the CRR, report as <

If a value exceeds the upper end of the CRR, report as >

Dilution protocol – none

PROCEDURE LIMITATIONS

≥2+ Hemolysis may yield a false elevation of the result.

For a complete listing of interferences, refer to the manufacturer’s package insert. If interferences are present in the sample and could affect results, the sample may need to be rejected. Consult with lab lead, lab director, or medical director for further guidance as needed.

REFERENCES

ABBOTT AST package insert