AMMONIA PROCEDURE FOR ALINITY

PRINCIPLE

Ammonia, derived from the catabolism of amino acids and from the action of intestinal bacteria on dietary protein, is converted to urea in the liver and rendered non-toxic. However, excess ammonia can have a toxic effect on the central nervous system, and may be observed in severe liver failure, genetic defects of the urea cycle, and other hereditary disorders.

Methodology: Glutamate Dehydrogenase

SCOPE

Ammonia Ultra is intended for the quantitative enzymatic determination of ammonia in plasma on the Alinity C analyzer performed in the Texas Health Allen laboratory.

SAFETY

Bloodborne Pathogen Risk – high

Chemical Risk – low

PPE Required – lab coats, gloves, splash shields (goggles, face shields, counter-mounted shields)

- Do not use components beyond the expiration date.

- Do not mix materials from different kit lot numbers.

- This product requires the handling of human specimens. It is recommended that all human sourced materials are considered potentially infectious and be handled in accordance with OSHA Standard on Bloodborne Pathogens.

SPECIMEN REQUIRMENTS

Plasma is the only acceptable specimen.

Use plasma in lithium heparin tubes and separate from red blood cells as soon after collection as possible.

Do NOT use hemolyzed specimens, as red blood cells contain ammonia levels 3 times that of plasma.

COLLECTION AND HANDLING

Contamination may be caused by cigarette smoke one the patient or the phlebotomist.

Sample Collection: Use standard phlebotomy procedures for Texas Health Allen. The collection tube must be completely filled and placed IMMEDIATELY on ice.  REJECT/RECOLLECT specimen if not placed on ice immediately. If a recollect cannot be obtained footnote the result with “Ammonia sample not placed on ice. Ammonia may be falsely elevated.

Note: The manufacturer’s recommendation is no more than 15 minutes from collection to the start of centrifiugation. Timing is especially critical for patients with liver disease. Delays over 15 minutes have been shown to increase ammonia levels even at 0C. The plasma must be immediately removed from the RBCs immediately after centrifugation.

SPECIMEN STORAGE STABILITY

20-25C no stability

2-8C is good for 2 hours

-20C is good for 3 weeks (if collected properly, centrifuged and plasma separated immediately and placed in the freezer.)

Note: Allow specimens to reach room temperature prior to testing. Stored specimens must be mixed prior to testing and recentrifuged if they contain particulate matter.

REAGENT HANDLING

- Upon receipt, place reagent cartridge in an upright position for at least 1 hour before use to allow bubbles that may have formed to dissipate.

- If a reagent cartridge is dropped, place in an upright position for 1 hour before use.

- Reagents are susceptible to the formation of foam and bubbles. Bubbles may interfere with the detection of the reagent level in the cartridge and cause insufficient reagent aspiration that may adversely affect results.

REAGENT STORAGE

- Unopened stored at 2-8C are good until the original expiration date.

- Opened reagent is stable until the expiration date if capped tightly and stored at 2-8C.

- Onboard reagent stability is 15 days.

Note: Reagent on board stability is determined by the instrument when the reagent is put into use and stored electronically. The instrument flags the user with information about expiration dates and will not perform testing beyond the reagent’s on board expiration date.

No reagents will be used beyond the expiration date.

MAINTENANCE

Perform daily or as needed maintenance. Refer to Alinity System Operations Manual. Daily, weekly, monthly, and quarterly maintenance operations and instructions are described in the online manual.

CALIBRATION

Calibration is stable for approximately 24 hours and calibration is required with each new lot number of reagent. Verify calibration with at least two levels of controls. If control results are not within acceptable ranges, recalibration may be necessary.

To order a calibration select CREATE ORDER from the home screen, then select CALIBRATION from the tabs along the top. Enter the rack number in the rack field and the position number in the position field. In the Assays section, select the assays to be ordered and select ADD ORDER tab along the bottom of the screen.

Calibration may be necessary in the following conditions:

- Out of range control values

- Significant shift or trend in Quality Control (QC) values

- A bulk solution is placed into service

- Any dispense system component is replaced

- Any c-side maintenance procedure is performed

CALIBRATION VERIFICATION

Calibration is validated by running 2 levels of QC after each calibration. If QC values are acceptable, the calibration is considered valid, and patients can be run.

QUALITY CONTROL

Material: BioRad Ethanol/Ammonia Control, levels 1 and 2.

Storage and Stability: Stable until the expiration date printed on the bottle when stored at -20C. Swirl gently to ensure homogeneity. Bring control to room temperature prior to use.

Frequency: 2 levels of QC are to be run every 24 hours.

Interpretation and Troubleshooting: If QC results do not meet the acceptance criteria, patient values may be suspect. Review QC results and acceptance criteria following a change of reagent or calibrator lot.

PROCEDURE

- Ensure that sufficient reagent is on board and that it is calibrated.

- Verify sample integrity, sufficient volume, and approved sample type.

- Load sample into appropriate rack and lane on the Alinity.

RESULT REPORTING

AMR:

CRR:

Critical:

If a value exceeds the low end of the CRR, report as <

If a value exceeds the upper end of the CRR, report as >

Dilution protocol – none

PROCEDURE LIMITATIONS

Hemolyzed samples cannot be used. Reject a hemolysis of 1+ or higher and notify the caregiver documenting your call in LIS.

Icteric samples may give an error code. Other sample interferences may give codes as well and not produce a result. In many cases, a dilution will dilute out the interference and a numerical result can be obtained. If the error code persists on the dilution, the sample is resulted as “unable to analyze due to interference” and the caregiver notified, and the call is documented. The sample can also be sent to a reference lab whose methodology is different and can obtain an accurate result if the physician request it be done.

Platelet contamination may cause falsely elevated results. Ensure the sample is centrifuged according to manufacturer’s instructions.

Reject the sample if it is not received on ice. The chilling of the sample reduces the RBC’s release of ammonia, so samples not placed immediately on ice will have falsely elevated ammonia levels.

Icterus (4+) may yield inaccurate results. Footnote according to HIL procedure.

Lipemia (2-4+) may yield inaccurate results. Footnote according to HIL procedure.

For a complete listing of interferences, refer to the manufacturer’s package insert. If interferences are present in the sample and could affect results, the sample may need to be rejected. Consult with lab lead, lab director, or medical director for further guidance as needed.

REFERENCES

ABBOTT Multigent Ammonia Ultra package insert