CHOLESTEROL PROCEDURE FOR ALINITY

PRINCIPLE

Cholesterol is measured to serve as an indicator of liver and kidney function, intestinal adsorption, coronary artery disease risk, and thyroid function. Cholesterol levels are important in the diagnosis and classification of hyperlipoproteinemias.

It is recommended by the Adult Treatment Panel of the National Cholesterol Education Program (NCEP) that all adults over 20 years of age should have a fasting lipoprotein profile (total cholesterol, LDL cholesterol, HDL cholesterol, and triglycerides) once every 5 years to screen for CAD.

Methodology: Enzymatic

SCOPE

The Cholesterol assay is used for the quantitation of cholesterol in human serum or plasma on the Alinity C analyzer performed by the laboratory at Texas Health Allen.

SAFETY

Bloodborne Pathogen Risk – high

Chemical Risk – low

PPE Required – lab coats, gloves, splash shields (goggles, face shields, counter-mounted shields)

- Do not use components beyond the expiration date.

- Do not mix materials from different kit lot numbers.

- This product requires the handling of human specimens. It is recommended that all human sourced materials are considered potentially infectious and be handled in accordance with OSHA Standard on Bloodborne Pathogens.

SPECIMEN REQUIRMENTS

Serum and plasma are acceptable specimens. Fasting specimens are recommended.

Serum: Use serum with or without gel barrier collected by standard venipuncture techniques in plastic tubes. Ensure complete clot formation has taken place prior to centrifugation. Separate from red blood cells as soon after collection as possible.

Plasma: Use plasma with or without gel barrier. Acceptable anticoagulants: lithium heparin. Separate from red blood cells as soon after collection as possible.

20-25C is good for 7 days

2-8C is good for 7 days

-20C is good for 3 months

Note: Allow specimens to reach room temperature prior to testing. Stored specimens must be mixed prior to testing and recentrifuged if they contain particulate matter.

REAGENT HANDLING

- Upon receipt, place reagent cartridge in an upright position before use to allow bubbles that may have formed to dissipate.

- If a reagent cartridge is dropped, place in an upright position for 1 hour before use.

- Do not invert reagent cartridges prior to use.

- Reagents are susceptible to the formation of foam and bubbles. Bubbles may interfere with the detection of the reagent level in the cartridge and cause insufficient reagent aspiration that may adversely affect results.

REAGENT STORAGE

- The unopened reagents are stable until the expiration date when stored at 2-8C.

- Reagent stability is 30 days if the reagent is uncapped and on-board.

Note: Reagent on board stability is determined by the instrument when the reagent is put into use and stored electronically. The instrument flags the user with information about expiration dates and will not perform testing beyond the reagent’s on-board expiration date.

- Calibrator unopened is good until the expiration date printed on the label when stored at 2-8C.

- Calibrator opened is stable for 7 days at 2-8C if kept tightly capped.

- Calibrator on-board stability is 5 days.

No reagents will be used beyond the expiration date.

MAINTENANCE

Perform daily or as needed maintenance. Refer to Alinity System Operations Manual. Daily, weekly, monthly, and quarterly maintenance operations and instructions are described in the online manual.

CALIBRATION

Calibration is stable for approximately 30 days and calibration is required with each new lot number of reagent. Verify calibration with at least two levels of controls. If control results are not within acceptable ranges, recalibration may be necessary.

To order a calibration select CREATE ORDER from the home screen, then select CALIBRATION from the tabs along the top. Enter the rack number in the rack field and the position number in the position field. In the Assays section, select the assays to be ordered and select ADD ORDER tab along the bottom of the screen.

Calibration may be necessary in the following conditions:

- Out of range control values

- Significant shift or trend in Quality Control (QC) values

- A bulk solution is placed into service

- Any dispense system component is replaced

- Any c-side maintenance procedure is performed

CALIBRATION VERIFICATION

Calibration is validated by running 2 levels of QC after each calibration. If QC values are acceptable, the calibration is considered valid and patients can be run.

QUALITY CONTROL

Material: BioRad Unassayed Chemistry Control, levels 1 and 2.

Storage and Stability: Unassayed Chemistry Control - Stable until the expiration date printed on the bottle when stored at -20C. Once thawed, the control is stable for 15 days at 2-8C, except for bilirubin for 6 days at 2-8C. Thaw the control bottle at room temperature (18-25C) until completely thawed. DO NOT THAW CONTROLS IN THE FRIDGE. Swirl gently to ensure homogeneity. Use immediately or refrigerate.

Frequency: 2 levels of Unassayed QC are to be run every 24 hours.

Interpretation and Troubleshooting: If QC results do not meet the acceptance criteria, patient values may be suspect. Review QC results and acceptance criteria following a change of reagent or calibrator lot.

PROCEDURE

- Ensure that sufficient reagent is on board and that it is calibrated.

- Verify sample integrity, sufficient volume, and approved sample type.

- Load sample into appropriate rack and lane on the Alinity.

RESULT REPORTING

AMR:

CRR:

Critical:

If a value exceeds the low end of the CRR, report as <

If a value exceeds the upper end of the CRR, report as >

Dilution protocol – none

PROCEDURE LIMITATIONS

Specimens for cholesterol displaying a 4+ hemolysis may yield falsely increased results by up to 12%.

Specimens for cholesterol displaying a 2+ or greater lipemia may yield falsely increased results by 13%.

For a complete listing of interferences, refer to the manufacturer’s package insert. If interferences are present in the sample and could affect results, the sample may need to be rejected. Consult with lab lead, lab director, or medical director for further guidance as needed.

REFERENCES

ABBOTT cholesterol package insert

ABBOTT multiconstituent Calibrator Package Insert