DIRECT BILIRUBIN PROCEDURE FOR ALINITY

PRINCIPLE

Red blood cells (RBCs) are broken down in the reticuloendothelial system at the end of their circulating life. The resulting heme, once the iron is removed, is then converted to bilirubin. This process accounts for about 80% of the bilirubin formed each day.

Once formed, bilirubin is transported to the liver bound to albumin as it is water insoluble. This fraction of bilirubin is referred to as indirect or unconjugated bilirubin. Bilirubin is conjugated in the liver and is now referred to as direct bilirubin. Direct bilirubin is excreted via the biliary system into the intestine. In normal circumstances elimination of bilirubin is almost complete and serum levels are negligible.

Direct bilirubin is elevated in conditions causing hepatic obstruction, hepatitis, cirrhosis, several inherited enzyme deficiencies, and inherited defects in canalicular excretion.

Methodology: Diazo Reaction

SCOPE

The Direct Bilirubin assay is used for the quantitation of direct bilirubin in human serum or plasma on the Alinity C analyzer performed in the Texas Health Allen laboratory.

SAFETY

Bloodborne Pathogen Risk – high

Chemical Risk – low

PPE Required – lab coats, gloves, splash shields (goggles, face shields, counter-mounted shields)

- Do not use components beyond the expiration date.

- Do not mix materials from different kit lot numbers.

- This product requires the handling of human specimens. It is recommended that all human sourced materials are considered potentially infectious and be handled in accordance with OSHA Standard on Bloodborne Pathogens.

SPECIMEN REQUIRMENTS

Serum and plasma without visible hemolysis or lipemia are acceptable specimens.

Serum: Use serum with or without gel barrier collected by standard venipuncture techniques in plastic tubes. Ensure complete clot formation has taken place prior to centrifugation. Separate from red blood cells as soon after collection as possible.

Plasma: Use plasma with or without gel barrier. Separate from red blood cells as soon after collection as possible.

SPECIMEN STORAGE STABILITY

Specimens should be protected from bright light as bilirubin is photolabile. Direct bilirubin is stable in serum and plasma as follows:

20-25C is good for 2 days

2-8C is good for 7 days

-20C is good for 3 months

Note: Allow specimens to reach room temperature prior to testing. Stored specimens must be mixed prior to testing and recentrifuged if they contain particulate matter.

REAGENT HANDLING

- Upon receipt, place reagent cartridge in an upright position for at least 1 hour before use to allow bubbles that may have formed to dissipate.

- If a reagent cartridge is dropped, place in an upright position for 1 hour before use.

- Reagents are susceptible to the formation of foam and bubbles. Bubbles may interfere with the detection of the reagent level in the cartridge and cause insufficient reagent aspiration that may adversely affect results.

REAGENT STORAGE

- Unopened stored at 2-8C are good until the original expiration date.

- Onboard reagent stability is 28 days.

Note: Reagent on board stability is determined by the instrument when the reagent is put into use and stored electronically. The instrument flags the user with information about expiration dates, and will not perform testing beyond the reagent’s on board expiration date.

No reagents will be used beyond the expiration date.

MAINTENANCE

Perform daily or as needed maintenance. Refer to Alinity System Operations Manual. Daily, weekly, monthly, and quarterly maintenance operations and instructions are described in the online manual.

CALIBRATION

Calibration is stable for approximately 14 days (336 hours) and calibration is required with each new lot number of reagent. Verify calibration with at least two levels of controls. If control results are not within acceptable ranges, recalibration may be necessary.

To order a calibration select CREATE ORDER from the home screen, then select CALIBRATION from the tabs along the top. Enter the rack number in the rack field and the position number in the position field. In the Assays section, select the assays to be ordered and select ADD ORDER tab along the bottom of the screen.

Calibration may be necessary in the following conditions:

- Out of range control values

- Significant shift or trend in Quality Control (QC) values

- A bulk solution is placed into service

- Any dispense system component is replaced

- Any c-side maintenance procedure is performed

CALIBRATION VERIFICATION

Calibration is validated by running 2 levels of QC after each calibration. If QC values are acceptable, the calibration is considered valid and patients can be run.

QUALITY CONTROL

Material: BioRad Unassayed Chemistry Control, levels 1 and 2. Pediatric Control, level 2.

Storage and Stability: Unassayed Chemistry Control - Stable until the expiration date printed on the bottle when stored at -20C. Once thawed, the control is stable for bilirubin for 6 days at 2-8C. Thaw the control bottle at room temperature (18-25C) until completely thawed. DO NOT THAW CONTROLS IN THE FRIDGE. Swirl gently to ensure homogeneity. Use immediately or refrigerate.

Storage and Stability: Pediatric Control – Stable until the expiration date on the bottle when stored at -20C. Once thawed and/or opened, the control is stable for 14 days at 2-8C. Thaw control bottle at room temperature until completely thawed. Swirl gently to ensure homogeneity. Use immediately or refrigerate.

Frequency: 2 levels of QC are to be run every 24 hours. One level of pediatric control is run every 24 hours.

Interpretation and Troubleshooting: If QC results do not meet the acceptance criteria, patient values may be suspect. Review QC results and acceptance criteria following a change of reagent or calibrator lot.

PROCEDURE

- Ensure that sufficient reagent is on board and that it is calibrated.

- Verify sample integrity, sufficient volume, and approved sample type.

- Load sample into appropriate rack and lane on the Alinity.

RESULT REPORTING

AMR:

CRR:

Critical:

If a value exceeds the low end of the CRR, report as <

If a value exceeds the upper end of the CRR, report as >

Dilution protocol – none

PROCEDURE LIMITATIONS

Some specimens may give a direct bilirubin result slightly greater than the total bilirubin result. During internal testing at Abbott Laboratories, specimens with total bilirubin concentrations at 0.2 mg/dL (3.4 umol/L) or less occasionally gave a direct bilirubin result that slightly exceeded their respective total bilirubin result. This may be observed when nearly all reacting bilirubin is direct bilirubin.

Specimens with a hemolysis of 3+ or greater, or specimens with triglyceride levels above 1000mg/dL may yield erroneous results.

Specimens containing indocyanine green may yield falsely elevated results.

≥2+ hemolysis may yield a decrease in the result by as much as 0.2 mg/dL.

≥2+ lipemia may yield a false elevation of the result by as much as 0.1 to 0.3 mg/dL.

For a complete listing of interferences, refer to the manufacturer’s package insert. If interferences are present in the sample and could affect results, the sample may need to be rejected. Consult with lab lead, lab director, or medical director for further guidance as needed.

REFERENCES

ABBOTT Direct Bilirubin package insert

ABBOTT Bilirubin Calibrator Package Insert