TOTAL PROTEIN PROCEDURE FOR ALINITY

PRINCIPLE

Plasma proteins derive primarily from synthesis in the liver, plasma cells, lymph nodes, spleen and bone marrow. In disease states both the total plasma protein level and the ratio of the individual fractions may be altered.

Hypoproteinemia (abnormally low levels of plasma protein) is observed in nephrotic syndrome, extensive bleeding, sprue (deficient protein absorption), severe burns, salt retention syndromes, and Kwashiorkor (acute protein starvation).

Hyperproteinemia (abnormally high levels of plasma protein) is observed in severe dehydration and multiple myeloma.

Changes in the proportions of the plasma proteins can occur in several of the protein fractions and often without alterations in the quantity of total protein. The A/G ratio has commonly been used as an index of the distribution between the albumin and globulin fractions. This ratio can be significantly altered in such conditions as cirrhosis of the liver, glomerulonephritis, nephrotic syndrome, acute hepatitis, lupus erythematosus, and some infections.

Methodology: Biuret

SCOPE

The Total Protein assay is used for the quantitation of total protein in human serum or plasma on the Alinity C analyzer performed in the Texas Health Allen laboratory.

SAFETY

Bloodborne Pathogen Risk – high

Chemical Risk – low

PPE Required – lab coat, gloves, splash guard (goggles, face shields, counter-mounted shields)

SPECIMEN REQUIREMENTS

Serum and Plasma are acceptable specimens.

Serum: use serum with or without gel barrier. Ensure complete clot formation has taken place prior to centrifugation. Separate from red blood cells as soon after collection as possible.

Plasma: use plasma (acceptable anticoagulants: lithium heparin) with or without gel barrier. Separate from red blood cells as soon after collection as possible.

SPECIMEN STABILITY

20-25C is good for 1 week

2-8C is good for 1 month

-20C is good for 2 months

REAGENT HANDLING

- Upon receipt, place the reagent cartridge in an upright position for at least 1 hour prior to use to allow bubbles that may have formed to dissipate.

- If a reagent cartridge is dropped, place in an upright position for 1 hour before use.

- Reagents are susceptible to form foam and bubbles. Bubbles may interfere with the detection of the reagent level in the cartridge and cause insufficient reagent aspiration that may adversely affect results.

REAGENT STORAGE

- Unopened reagent is stable until the expiration date when stored at 15-30C.

- Uncapped On-board reagent is stable for 23 days.

  - Reagent on board stability is determined by the instrument when the reagent is put into use and stored electronically.

  - The instrument flags the operator with information about expired reagents and will not perform testing beyond the reagent’s on-board expiration date.

CALIBRATOR STABILITY

- Unopened – until expiration date on bottle when stored at 2-8C

- Opened – 7 days at 2-8C

No reagents will be used beyond the expiration date.

MAINTENANCE

Perform daily or as needed maintenance. Refer to Alinity C System Operations Manual, maintenance section. Daily, weekly, monthly, and quarterly maintenance operations and instructions are described in the online manual.

CALIBRATION

Calibration is stable for approximately 23 days (552 hours) and calibration is required with each reagent lot number change.

Verify calibration with at least 2 levels of control. If control results are not acceptable, recalibration may be necessary.

To order a calibration select CREATE ORDER from the home screen, then select CALIBRATION from the tabs along the top. Enter the rack number in the rack field and the position number of the first calibrator spot in the position field. In the Assays section select the assay to be ordered then select ADD ORDER tab along the bottom of the screen. Load the calibrators by loading into the sample loading slots.

For information on calibrator values, refer to the Multiconstituent Calibrator (MCC) package insert. Calibration values are programmed in the instrument assay parameters at each calibrator lot number change. Calibrator value information can also be found in the assay configuration tab.

Calibration may be necessary in the following conditions:

- Out-of-range control values

- Significant shift or trend in QC values

- A bulk solution is placed into service

- Any c-side system component is replaced

- C-side maintenance is performed in response to instrument troubleshooting

CALIBRATION VERIFICATION

Calibration is validated by running 2 levels of control after each calibration. If QC values are acceptable according to established ranges, the calibration is considered valid, and patients can then be run.

QUALITY CONTROL

Material: Serum: BioRad Unassayed chemistry control levels 1 and 2.

Storage and Stability: stable until expiration date printed on the bottle when stored unopened at -20C. Once thawed, the control is stable for 15 days at 2-8C, except bilirubin which is stable for 6 days. Thaw control bottle at room temperature (18-25C) until completely thawed. Swirl gently to mix, use immediately or refrigerate.

Frequency: two levels of QC are to be run every 24 hours.

Interpretation and Troubleshooting: If QC results are not within range, patient values may be suspect. Review QC results and acceptance criteria following a change of reagent or calibrator lot.

PROCEDURE

For detailed instructions in Alinity C operations, refer to section 5 of the instrument manual, and the Alinity Operation SOP.

ASSAY PROCEDURE

- Ensure that sufficient reagent is on board and that it is calibrated.

- Verify sample integrity, sufficient volume, and approved sample type.

- Load sample into appropriate rack and lane on the Alinity

RESULT REPORTING

AMR:

CRR:

Critical:

If value exceeds the low end of the CRR, report as <

If a value exceeds the upper limit of the CRR, report as >

Dilution protocol – none

LIMITATIONS

- ≥3+ hemolysis may result in a false elevation of the result by as much as 12%.

- For a complete list of interferences, refer to the manufacturer’s package insert. If interferences are present in the sample and could affect results, the sample may have to be rejected. Consult with lab lead, lab director or medical director for further instruction as required.

REFERENCES

ABBOTT  Total Protein package insert

ABBOTT Multiconstituents Calibrator package insert