UREA NITROGEN PROCEDURE FOR ALINITY

PRINCIPLE

Blood Urea Nitrogen (BUN) measurements are used in the diagnosis of certain renal and metabolic diseases. The determination of serum urea nitrogen is a widely used test for the evaluation of kidney function. Used in conjunction with the serum creatinine test, the BUN can help differentiate prerenal, renal and postrenal diseases,

Prerenal Diseases: cardiac decompensation, water depletion, increased protein catabolism

Renal Diseases: glomerulonephritis, chronic nephritis, polycystic kidney, nephrosclerosis tubular necrosis

Postrenal Diseases: obstruction of the urinary tract

Methodology: Urease

SCOPE

The Urea Nitrogen assay is used for the quantitation of urea nitrogen in human serum or plasma on the Alinity C analyzer performed in the Texas Health Allen laboratory.

SAFETY

Bloodborne Pathogen Risk – high

Chemical Risk – low

PPE Required – lab coat, gloves, splash guard (goggles, face shields, counter-mounted shields)

SPECIMEN REQUIREMENTS

Serum and plasma are acceptable specimens. Patient status is preferably fasting.

Serum: Use serum with or without gel barrier. Ensure complete clot formation has taken place prior to centrifugation. If the specimen is centrifuged before a complete clot forms, the presence of fibrin may cause erroneous results. Separate from RBCs as soon after collections as possible.

Plasma: Use plasma with or without gel barrier including green microtainer tubes (acceptable anticoagulant: lithium heparin). Separate from RBCs as soon as possible after collection.

SPECIMEN STABILITY:

20-25C is good for 7 days

2-8C is good for 7 days

-20C is good for 1 year

REAGENT HANDLING

- Upon receipt, place the reagent cartridge in an upright position for at least 1 hour prior to use to allow bubbles that may have formed to dissipate.

- If a reagent cartridge is dropped, place in an upright position for 1 hour before use.

- Reagents are susceptible to form foam and bubbles. Bubbles may interfere with the detection of the reagent level in the cartridge and cause insufficient reagent aspiration that may adversely affect results.

REAGENT STORAGE

- Unopened reagent is stable until the expiration date when stored at 2-8C.

- Uncapped On-board reagent is stable for 25 days.

  - Reagent on board stability is determined by the instrument when the reagent is put into use and stored electronically.

  - The instrument flags the operator with information about expired reagents and will not perform testing beyond the reagent’s on-board expiration date.

CALIBRATOR STABILITY

- Unopened – until expiration date on bottle when stored at 2-8C

- Opened – 7 days at 2-8C

No reagents will be used beyond the expiration date.

MAINTENANCE

Perform daily or as needed maintenance. Refer to Alinity C System Operations Manual, maintenance section. Daily, weekly, monthly, and quarterly maintenance operations and instructions are described in the online manual.

CALIBRATION

Calibration is stable for approximately 7 days (168 hours) and calibration is required with each reagent lot number change.

Verify calibration with at least 2 levels of control. If control results are not acceptable, recalibration may be necessary.

To order a calibration select CREATE ORDER from the home screen, then select CALIBRATION from the tabs along the top. Enter the rack number in the rack field and the position number of the first calibrator spot in the position field. In the Assays section select the assay to be ordered then select ADD ORDER tab along the bottom of the screen. Load the calibrators by loading into the sample loading slots.

For information on calibrator values, refer to the Multiconstiituent Calibrator (MCC) package insert. Calibration values are programmed in the instrument assay parameters at each calibrator lot number change. Calibrator value information can also be found in the assay configuration tab.

Calibration may be necessary in the following conditions:

- Out-of-range control values

- Significant shift or trend in QC values

- A bulk solution is placed into service

- Any c-side system component is replaced

- C-side maintenance is performed in response to instrument troubleshooting

CALIBRATION VERIFICATION

Calibration is validated by running 2 levels of control after each calibration. If QC values are acceptable according to established ranges, the calibration is considered valid, and patients can then be run.

QUALITY CONTROL

Material: Serum: BioRad Unassayed chemistry control levels 1 and 2.

Storage and Stability: serum: stable until expiration date printed on the bottle when stored unopened at -20C. Once thawed, the control is stable for 15 days at 2-8C, except bilirubin which is stable for 6 days. Thaw control bottle at room temperature until completely thawed. Swirl gently to ensure homogeneity. Use immediately or refrigerate.

Frequency: two levels of QC are to be run every 24 hours.

Interpretation and Troubleshooting: If QC results are not within range, patient values may be suspect. Review QC results and acceptance criteria following a change of reagent or calibrator lot.

PROCEDURE

For detailed instructions in Alinity C operations, refer to section 5 of the instrument manual, and the Alinity Operation SOP.

ASSAY PROCEDURE

- Ensure that sufficient reagent is on board and that it is calibrated.

- Verify sample integrity, sufficient volume, and approved sample type.

- Load sample into appropriate rack and lane on the Alinity

RESULT REPORTING

AMR:

CRR:

Critical:

If value exceeds the low end of the CRR, report as <

If a value exceeds the upper limit of the CRR, report as >

Dilution protocol – none

LIMITATIONS

- No significant interference from hemolysis, icterus, or lipemia are listed for serum and plasma.

- For a complete list of interferences, refer to the manufacturer’s package insert. If interferences are present in the sample and could affect results, the sample may have to be rejected. Consult with lab lead, lab director or medical director for further instruction as required.

REFERENCES

ABBOTT Urea Nitrogen package insert

ABBOTT Multiconstituent Calibrator package insert