URINE/CSF TOTAL PROTEIN PROCEDURE FOR ALINITY

PRINCIPLE

Urine: The role of the renal system in the conservation of plasma proteins has been recognized for some time. Under normal physiological conditions small molecular weight proteins, such as insulin, pass through the glomeruli in relatively large amounts. Intermediate sized proteins, such as transferrin and albumin, also pass through but in relatively small amounts. Most of these proteins are reabsorbed in the renal tubules.

CSF (Cerebrospinal fluid): Most CSF protein originates by diffusion from plasma across the blood-CSF barrier. Increased levels occur because of increased permeability of the blood-CSF barrier.

Methodology: Benzethonium Chloride

SCOPE

The Urine/CSF Total Protein assay is used for the quantitation of protein in human urine or CSF on the Alinity C analyzer performed in the Texas Health Allen laboratory. CSF measurements aid in the diagnosis and treatment of: meningitis, brain tumors, and infections of the central nervous system.

SAFETY

Bloodborne Pathogen Risk – high

Chemical Risk – low

PPE Required – lab coat, gloves, splash guard (goggles, face shields, counter-mounted shields)

SPECIMEN REQUIREMENTS

Urine and CSF are acceptable specimens.

Urine: 24 hour timed urine specimens are preferred, with no preservative.

Both urine and CSF are to be centrifuged prior to testing.

SPECIMEN STABILITY

Urine:

20-25C is good for 24 hours

2-8C is good for 7 days

-20C is good for 1 month

CSF:

20-25C is good for 24 hours

2-8C is good for 7 days

-20C is good for 1 month

REAGENT HANDLING

- Upon receipt, place the reagent cartridge in an upright position for at least 1 hour prior to use to allow bubbles that may have formed to dissipate.

- If a reagent cartridge is dropped, place in an upright position for 1 hour before use.

- Reagents are susceptible to form foam and bubbles. Bubbles may interfere with the detection of the reagent level in the cartridge and cause insufficient reagent aspiration that may adversely affect results.

REAGENT STORAGE

- Unopened reagent is stable until the expiration date when stored at 15-30C.

- Uncapped On-board reagent is stable for 41 days.

  - Reagent on board stability is determined by the instrument when the reagent is put into use and stored electronically.

  - The instrument flags the operator with information about expired reagents and will not perform testing beyond the reagent’s on-board expiration date.

CALIBRATOR STABILITY

- Unopened – until expiration date on bottle when stored at 2-8C

- Opened – 41 days at 2-8C

No reagents will be used beyond the expiration date.

MAINTENANCE

Perform daily or as needed maintenance. Refer to Alinity C System Operations Manual, maintenance section. Daily, weekly, monthly, and quarterly maintenance operations and instructions are described in the online manual.

CALIBRATION

Calibration is stable for approximately 41 days (984 hours) and calibration is required with each reagent lot number change.

Verify calibration with at least 2 levels of control. If control results are not acceptable, recalibration may be necessary.

To order a calibration select CREATE ORDER from the home screen, then select CALIBRATION from the tabs along the top. Enter the rack number in the rack field and the position number of the first calibrator spot in the position field. In the Assays section select the assay to be ordered then select ADD ORDER tab along the bottom of the screen. Load the calibrators by loading into the sample loading slots.

For information on calibrator values, refer to the Urine/CSF Protein Calibrator package insert. Calibration values are programmed in the instrument assay parameters at each calibrator lot number change. Calibrator value information can also be found in the assay configuration tab.

Calibration may be necessary in the following conditions:

- Out-of-range control values

- Significant shift or trend in QC values

- A bulk solution is placed into service

- Any c-side system component is replaced

- C-side maintenance is performed in response to instrument troubleshooting

CALIBRATION VERIFICATION

Calibration is validated by running 2 levels of control after each calibration. If QC values are acceptable according to established ranges, the calibration is considered valid, and patients can then be run.

QUALITY CONTROL

Material: Serum: BioRad CSF control levels 1 and 2.

Storage and Stability: CSF control - stable until expiration date printed on the bottle when stored unopened at 2-8C. Once thawed, the control is stable for 30 days at 2-8C.

Urine control – stable until expiration date printed on bottle when stored unopened at 2-8C. Once opened, the stability is 30 days at 2-8C.

Frequency: two levels of QC are to be run every 24 hours.

Interpretation and Troubleshooting: If QC results are not within range, patient values may be suspect. Review QC results and acceptance criteria following a change of reagent or calibrator lot.

PROCEDURE

For detailed instructions in Alinity C operations, refer to section 5 of the instrument manual, and the Alinity Operation SOP.

ASSAY PROCEDURE

- Ensure that sufficient reagent is on board and that it is calibrated.

- Verify sample integrity, sufficient volume, and approved sample type.

- Load sample into appropriate rack and lane on the Alinity

RESULT REPORTING

AMR:

CRR:

Critical:

If value exceeds the low end of the CRR, report as <

If a value exceeds the upper limit of the CRR, report as >

Dilution protocol – none

LIMITATIONS

- For a complete list of interferences, refer to the manufacturer’s package insert. If interferences are present in the sample and could affect results, the sample may have to be rejected. Consult with lab lead, lab director or medical director for further instruction as required.

REFERENCES

ABBOTT U-CSF protein package insert

ABBOTT U-CSF Calibrator package insert